## zResults
Distribution | Lab Code | Analytical Technique | 206Pb/238U Age Result | 206Pb/238U Age Uncertainty | 207Pb/206Pb Age Result | 207Pb/206Pb Age Uncertainty | 208Pb/232Thb Age Result | 208Pb/232Th Age Uncertainty | Total Measurements | Total Measurements in final pool | Total Range in final pool | Common Pb Correction | ID / Which type of instrument did you use? | ID / Which spike did you use? | ID / Did you mechanically abrade your test portions? | ID / Did you chemically abrade your test portions? | ID / What range of masses (mg) of test portions did you use? From value | ID / What range of masses (mg) of test portions did you use? To value | ID / What range in 206/204 did you observe? | ID / How many hours did you use for digesting this material? | ID / What temperature (°C) did you use for digesting this material? | ID / What acid did you use to digest the test portions? | ID / Did you correct for thorium disequilibrium? | ID / What 238U/235U ratio did you apply? | ID / What 238U decay constant (y-1) did you use? | ID / What 235U decay constant (y-1) did you use? | ID / Which data reduction package did you use? | ID / Are your reported ages based on simple or weighted means? |  |  | Other Relevant Information | Comments
G-Chron 2019 | A57 | ID | 295.529 | 0.38 | 295.46 | 0.47 |  |  | 16 | 6 | 294.59 - 297.42 | 204Pb ancient | TIMS | ET535 |  |  | 10 | 100 | 1200 - 17000 | 72 + 48 in two stages | 220 | Concentrated HF admixed with HNO3 | Yes | 137.88 | 1.5513e-10 | 9.8485e-10 | 2006 | Weighted Mean |  |  | Mass before chemical abrasion; accurate mass not possible. Dissolution-pure HF and HNO3 in spike. | Data are concordant and precise, but spread up-and-down within the band of decay constant uncertainty; excess scatter present in mean age calculations. Ten analyses were rejected due to unacceptably large Pb blanks.
Distribution | Lab Code | Analytical Technique | 206Pb/238U Age Result | 206Pb/238U Age Uncertainty | 207Pb/206Pb Age Result | 207Pb/206Pb Age Uncertainty | 208Pb/232Th Age Result | 208Pb/232Th Age Uncertainty | Total Measurements | Total Measurements in final pool | Total Range in final pool | Common Pb Correction | SIMS / Which type of instrument did you use? | SIMS / What type of primary ion source did you use? | SIMS / What type of primary species did you use? | SIMS / Did you use oxygen flooding? | SIMS / Which primary mode did you use? | SIMS / What primary ion current did you use? | SIMS / What beam diameter did you use? | SIMS / Which calibration material did you use? | SIMS / What was the 206Pb/238U age (Ma) that you assigned to your calibration material? | SIMS / What was the 206Pb/238U standard error (Ma) of the mean on your calibration material? | SIMS / How many analyses were conducted on your calibration material? | SIMS / How many determinations on your calibration material were rejected as outliers? | SIMS / Which calibration scheme did you use? | SIMS / Which curve fitting scheme did you use? |  |  |  |  | Other Relevant Information | Comments
G-Chron 2019 | A19 | SIMS | 295.5 | 0.7 | 277 | 18.5 |  |  | 23 | 22 | 288.7-301.3 | other | SHRIMP | Duoplasmatron | 16 O - | No | Kohler | > 5 to 10 nA | > 10 to 20 µm | Temora2 | 416.8 |  | 14 | 0 | 206/238 vs 254/238 | Power law |  |  |  |  | IP966_2, 1sigma spot to spot error = 0.68%, 0.8% used (our lab minimum), 1s error of the mean for this calibration 0.24%, 6266 used as concentration reference material. | One grain had very elevated common Pb (204Pb) counts as well as high background. No other grains demonstrated this feature. Replicate analyses of this grain were not conducted and this analysis was excluded from the mean. The error reported above only considers in the internal sources of error and thus would only be appropriate when comparing this result to other samples on the same mount. Our recognized long term reproducibility is 1% (2s) and thus would be 295.5±1.5 Ma (reported as 1s as per the G-Chron format)
G-Chron 2019 | A29 | SIMS | 296.2 | 0.5 | 287.6 | 6.6 |  |  | 90 | 89 | 13 | 204Pb modern | Cameca 1270/1280/1300 | Duoplasmatron | 16 O 2- | Yes | Kohler | > 10 to 15 nA | > 20 to 30 µm | 91500 | 1062 | 3.1 | 19 | 0 | 206/238 vs 270/238 | Power law |  |  |  |  | The finial uncertainty for single sport of unknowns were propogated from internal error of measured Pb/U and the standard deviation of 91500 standard (0.5% was used, though the real is 0.44%). | With 91500 zircon (U-Pb=1062 Ma) as standard, Plesovice and TEMORA were measured together. The 20 measurements on 91500 yielded standard deviation of U-Pb age of 0.47% (1 SD), for 20 measurements on Plevosice is 0.62% (1 SD), for 20 measurements on TEMORA is 0.55% (1 SD). Ninety analyese, excluding one outlier at 285 Ma, yielded 0.8% (1 SD) for U-Pb age. The common lead level is very low, similar to other standards.
G-Chron 2019 | A47 | SIMS | 295.5 | 0.5 | 301 | 14 | 292.1 | 1.8 | 32 | 32 | 288.0 - 305.8 | 204Pb ancient | SHRIMP | Duoplasmatron | 16 O 2- | No | Kohler | > 2 to 5 nA | > 20 to 30 µm | Temora2 | 416.8 | 1.6 | 20 | 0 | 206/238 vs 254/238 | Linear |  |  |  |  | 91500 zircon as U content RM (n=3). Analytical parameters: secondary accelerating voltage 1keV, exit slit size 90 mkm, mass-resolution >6000 at 1% peak-high, Kohler focusing aperture size of 100 mkm | 30 fragments were conducted, their CL, BSE and optic images were obtained, 2 additional analyses were measured in two most uranium fragments. U content ranged from 37 upto 667 ppm, mean Th/U=0.67. Arens-Wetherill concordant age was 295.4 ± 0.52 (mswd=0.025). Zircons M257 and GJ1-59 were applied as secondary RM and yielded concordant (AW) ages of 560.1 ± 0.6 Ma (n=10) and 608.9 ± 0.9 Ma (n=9) correspondingly.
G-Chron 2019 | A55 | SIMS | 292.4 | 1.15 | 301 | 38 |  |  | 38 | 37 | 282-307 | other | SHRIMP | Duoplasmatron | 16 O 2- | No | Kohler | <= 2 nA | > 20 to 30 µm | BR266 | 559.1 Ma | 0.25% | 20 | 1 | 206/238 vs 254/238 | Logarithmic | [5 empty] | Four large chips were used. Highly variable Th/U (0.5-1.5) and U concentration (34-440 µg/g).
G-Chron 2019 | A64 | SIMS | 295.7 | 3.3 | 301 | 36 |  |  | 25 | 25 | 11.7 | 204Pb modern | Cameca 1270/1280/1300 | Duoplasmatron | 16 O 2- | Yes | Kohler | > 10 to 15 nA | > 20 to 30 µm | 91500 | 1062.4 | 1.9 Ma | 16 | 0 | 206/238 vs 270/238 | Power law |  |  |  |  | Mass resolution of about 5400 was achieved with entrance slit of 80 um, exit slit of 173 um, energy slit of 50 eV and field aperture of 6000. | Measured 206Pb/204Pb ratios range from 4205 to 43991. Reported uncertainties for the mean 206Pb/238U age and 207Pb/206Pb age are 1sd level.
G-Chron 2019 | A76 | SIMS | 296.6 | 0.8 | 295 | 14.5 | 291 | 1.7 | 20 | 20 | 7.9 | 204Pb modern | Cameca 1270/1280/1300 | RF plasma source | 16 O 2- | Yes | Gaussian with raster | > 10 to 15 nA | > 10 to 20 µm | 91500 | 1062.4 Ma | 8.5 Ma | 14 | 1 | 206/254 vs 270/254 | Power law |  |  |  |  | Mono-coll. into ETP EM. Sec. "XY" mode. Impact energy 23 kV; energy window 45 eV. Transfer magn. ca. 160x = ~18 um f.o.v.). Presputter 25 um for 60 sec.; in-run raster 5 um. # of cycles = 8. | There is a wide range in U concentration from ca. 50 to 200 ppm. Thus, and given the young age, 207Pb is highly susceptible to overcorrection for 204 (204 corrected 207Pb/206Pb in this dataset was ca. 261 Ma.
Distribution | Lab Code | Analytical Technique | 206Pb/238U Age Result | 206Pb/238U Age Uncertainty | 207Pb/206Pb Age Result | 207Pb/206Pb Age Uncertainty | 208Pb/232Th Age Result | 208Pb/232Th Age Uncertainty | Total Measurements | Total Measurements in final pool | Total Range in final pool | Common Pb Correction | LA-ICP-MS: QuadToF / Which type of instrument did you use? | LA-ICP-MS: QuadToF / Please specify the make and model of the mass spectrometer that you used | LA-ICP-MS: QuadToF / Which laser wavelength did you use? | LA-ICP-MS: QuadToF / Which laser pulse length did you use? | LA-ICP-MS: QuadToF / Which laser fluence did you use? | LA-ICP-MS: QuadToF / What laser spot diameter did you use? | LA-ICP-MS: QuadToF / What sampling mode did you use? | LA-ICP-MS: QuadToF / Sampling mode - enter scan speed (μm/s) if applicable | LA-ICP-MS: QuadToF / Which carrier gas did you use? | LA-ICP-MS: QuadToF / Did you use a collision cell? | LA-ICP-MS: QuadToF / Did you use a reaction cell? | LA-ICP-MS: QuadToF / What laser repletion rate (Hz) did you use? | LA-ICP-MS: QuadToF / Did you use a signal smoothing device? | LA-ICP-MS: QuadToF / Which calibration material did you use? | LA-ICP-MS: QuadToF / What was the 206Pb/238U age (Ma) that you assigned to your calibration material? | LA-ICP-MS: QuadToF / What was the 206Pb/238U standard error (Ma) of the mean on your calibration material? | LA-ICP-MS: QuadToF / How many analyses were conducted on your calibration material? | LA-ICP-MS: QuadToF / How many determinations on your calibration material were rejected as outliers? | Other Relevant Information | Comments
G-Chron 2019 | A1 | LA-ICP-MS: QuadToF | 294.8 | 0.7 | 299 | 16 | 255.4 | 2.6 | 15 | 13 | 298.6-290.8 | none | Single quadrupole | Bruker M90 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 6 | No | GJ-1 | 600 | 3 | 6 | 0 |  | two points have very low Pb
G-Chron 2019 | A3 | LA-ICP-MS: QuadToF | 293.4 | 1.6 | 296 | 42 |  |  | 35 | 32 | 299.9-287.9 | none | Single quadrupole | Agilent7500Ce | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | No | 91500 | 1062.7 | 6.7 | 18 | 1
G-Chron 2019 | A4 | LA-ICP-MS: QuadToF | 293.79 | 0.97 | 301 | 26 | 295.9 | 5.3 | 63 | 62 | 8.93 | none | Single quadrupole | Analytikjena PlasmaQuant MS | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 Hz | No | GJ-1 | 600.3 | 3.1 | 16 | 0
G-Chron 2019 | A5 | LA-ICP-MS: QuadToF | 292.3 | 0.85 | 277 | 15 |  |  | 20 | 19 | 305.1-284.3 | none | Single quadrupole | Thermo iCAP-RQ | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | GJ-1 | 600 | 600 | 14 | 0 | Beam size: 30 um, Fluence: 2 J/cm2, Repeatition Rate: 5 Hz | At 30 um beam size, 207Pb signal intensity is low.
G-Chron 2019 | A6 | LA-ICP-MS: QuadToF | 294.76 | 1.59 | 355.4 | 30.2 | 293.78 | 2.47 | 30 | 30 | 280.05-309.6 | none | Triple quadrupole | Agilent 8900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 6 | Yes | 91500 | 1065 | 2.38 | 15 | 0 | R33 was run as a secondary reference material and was used to apply the uncertainty propagation protocol of Horstwood et al. (2016) | Our 7/6 data clearly are not good enough to get a meaningful 7/6 age by quadrupole ICP-MS on something this young, as I would expect.
G-Chron 2019 | A7 (*) | LA-ICP-MS: QuadToF | 301.2 | 5.2 | 284.4 | 32.6 |  |  | 16 | 13 | 21.8 | none | Single quadrupole | iCAP-Q | Solid state 213 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No |  | No | 91500 | 1062.4 | 1.9 | 18 | 0 | Pb blank was a bit high and sensitivity a bit low, so 207Pb/206Pb is not our best work. Sorry! | Unfortunately my Pb blank was a bit high and sensitivity a bit lower, so my 207Pb/206Pb ages are not very precise. I think I could do better with improved 207Pb measurements! Thank you!
G-Chron 2019 | A8 | LA-ICP-MS: QuadToF | 291 | 1.4 | 311 | 18 | 294 | 3.1 | 71 | 69 | 302.5 - 280.6 | none | Single quadrupole | Agilent 7900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 11 | Yes | 91500 | 1062.4 Ma | 4.1 Ma (2 standard errors) | 33 | 1 | 250 shots, 3 secondaries (GZ7, Plesovice, WRS-1348) also per standard block. 11 seconds between analyses. | Seems pretty good!
G-Chron 2019 | A10 | LA-ICP-MS: QuadToF | 294.1 | 2 | 235 | 45.1 |  |  | 30 | 1 | 302.5-281.0 | none | Single quadrupole | Thermo X-series II | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | Yes | Temora2 | 417 | 2.9 | 14 | 0 | He+Ar+N to MS | Relatively low U sample, so used Timora2 as the calibration std. Quality control vis Plec + Fc1
G-Chron 2019 | A12 | LA-ICP-MS: QuadToF | 293.92 | 0.37 | 288.1 | 7.4 |  |  | 35 | 35 | 298-288.3 | none | Single quadrupole | Agilent 7700x | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | GJ-1 | 601.8 | 0.4 | 15 | 0
G-Chron 2019 | A18 | LA-ICP-MS: QuadToF | 296.9 | 11.7 | 568.5 | 478.1 |  |  | 30 | 29 | 46.2 | none | Time-of-Flight | icpTOF | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | No |  | 523 | 5 | 4 | 0 | the inhouse stardard we used is SSEF87244, which is a yellow color single crystal zircon. Independent LA-ICP-MS has been measured on it to determine the age and confirm its homogeneity in age. | Please note that 207Pb/206Pb age [568.5 +- 478.1 (Avg. +- 1s)] can not be confidently determined by our method, this is due to the low intensity of 207Pb signal. (because of its low concentration, and small crater size) Therefore, we would NOT like to report this results in the submission, and will NOT use this result in routine analysis. We have done 6 round analyses of RAK17 ) in this G-Chron result, each round has 4 analysis on our inhouse standard. Therefore, the total number of analysis on our standard is 4 analyses * 6 rounds = 24 (no standard analyses have been rejected)
G-Chron 2019 | A21 | LA-ICP-MS: QuadToF | 298.7 | 2.5 | 292 | 25 | 290.6 | 4.8 | 22 | 22 | 290-311 | 204Pb modern | Single quadrupole | Agilent 7500cx | Solid state 213 nm | Nanosecond | > 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | Yes | No | 5Hz | No | GJ-1 | 608.5 | 0.4 | 4 | 2
G-Chron 2019 | A22 | LA-ICP-MS: QuadToF | 295 | 0.9 | 296.9 | 22.6 |  |  | 20 | 19 | 303.2-286.2 | other | Triple quadrupole | Agilent 8900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | Yes | No | 8 | Yes | 91500 | 1065 | 2.98 | 25 | 0 | Standard block is located in a 1" round independent ESI NWR 193 Laser. Collision cell gas is also He. Carrier gas flow is 650 mL/minute (optimized to minimize ThO to <0.2%) | Validation standard run simultaneously with RAK-17 were R33 (Mean 206Pb/238U Age = 416.1±1.0 (1S; n - 21) and Plesovice (Mean 206Pb/238U Age =336.0±0.9 (1S; n - 41). Plesovice was run in greater quantities to test for spatial variation within our laser chamber whereas R33 was mounted in the same epoxy puck as RAK-17. Our typical propagated uncertainty internal and systematic are typically 2-3% in the <redacted> lab. We have not yet published a paper where methods are described in detail. The RAK-17 ablation behavior was similar to R33 in many respects.
G-Chron 2019 | A23 | LA-ICP-MS: QuadToF | 298.1 | 2.5 | 297.1 | 14.8 | 290.1 | 6.9 | 65 | 63 | 278 - 314 | none | Triple quadrupole | Agilent 8800 QQQ | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 40 to 90 µm | Spot |  | Helium | No | No | 30 | No | 91500 | 1065 | 5.7 | 25 | 0
G-Chron 2019 | A24 | LA-ICP-MS: QuadToF | 294.1 | 1.1 | 248 | 25 |  |  | 15 | 12 | 9 Ma | none | Triple quadrupole | Agilent 8800 | Solid state 213 nm | Nanosecond | > 5 to 10 J cm-2 | > 40 to 90 µm | Spot |  | Helium | No | No | 10 | Yes | 91500 | 1062.4 | 2 | 5 | 0
G-Chron 2019 | A25 | LA-ICP-MS: QuadToF | 295.2 | 0.55 | 319 | 9 | 258.2 | 2 | 30 | 27 | 307-288 | none | Single quadrupole | Aglilent 7500cx | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | Yes | GJ-1 | 600 | 3 | 10 | 0 |  | we test the low U Th Pb zircon with white CL image
G-Chron 2019 | A28 | LA-ICP-MS: QuadToF | 283 | 5 | 407.9 | 110 |  |  | 30 | 22 | 266-293 | none | Single quadrupole | Agilent 7500ce | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | No | 91500 | 1065.4 | 6.4 | 16 | 5 | data reduction and calculations: PepiAGE and Isoplot | some grains after mounting in resin and polishing were very thin. As a consequence some laser spots went through grain and needed to be rejected. Concordia age was 282.5 ± 3.0 Ma.
G-Chron 2019 | A33 | LA-ICP-MS: QuadToF | 294.7 | 1.2 | 279 | 24 | 298.6 | 2 | 45 | 44 | 23 | none | Single quadrupole | Thermo Scientific XserieII | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | No | 91500 | 1065 | 8.6 | 8 | 0 | 2ml/min of N2 + 1L/min of Ar introduce to carrier gaz after ablation cell. quality control on plesovice ( as unknow) 338.8+-3.9Ma on 206/238 age n=4 | low Pb content estimated between 2 and 15ppm (U between 50 to 350 ppm) zonation observable on MEB image of polished section but no relevent diffrence for U/Pb age estimation.
G-Chron 2019 | A35 | LA-ICP-MS: QuadToF | 290.5 | 3.1 | 244 | 36 | 272 | 12 | 27 | 27 | 266 - 304.2 | none | Triple quadrupole | Perkin Elmer Nexion 2000 | Solid state 213 nm | Other | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | No | 91500 |  |  | 27 |  |  | We used 25micron crater size, 20 sec backround, 30 sec dwelltime and 20 sec washout time in measuments.
G-Chron 2019 | A36 | LA-ICP-MS: QuadToF | 293.21 | 0.86 | 0 | 0 |  |  | 45 | 35 | 14.07 | none | Triple quadrupole | Agilent 8900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 8 | Yes | GJ-1 | 600.4 | 0.174 | 28 | 2 | Data reduction done using GLITTER | Our method uses a small spot size and wide range of trace elements, so 7/6 ages for zircons < 1Ga are not normally reported. Analyses of some grains indicated the presence of inclusions containing Al, P and La
G-Chron 2019 | A38 | LA-ICP-MS: QuadToF | 296.29 | 1.47 | 296.04 | 4.21 |  |  | 32 | 30 | 287 - 305 | none | Single quadrupole | Agilent 7900x | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | 91500 | 1065 | 1.79 | 15 | 0
G-Chron 2019 | A39 | LA-ICP-MS: QuadToF | 294.2 | 0.5 | 224 | 15 |  |  | 30 | 29 | 9 | none | Triple quadrupole | Agilent 8900 QQQ | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Argon | No | No | 5 | Yes | 91500 | 1062.3 | 1.2 | 18 | 1 |  | Pulse-Analog limits on 8900 were not optimised for Pb so 207/206 ratios may be sub-optimal on this young material.
G-Chron 2019 | A40 | LA-ICP-MS: QuadToF | 289.6 | 3.4 | 281.8 | 35.4 | 289.2 | 5.6 | 14 | 13 | 11.3 | none | Single quadrupole | Agilent 7900 Quadropole | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 40 to 90 µm | Spot |  | Helium | No | No | 10 | Yes | 91500 | 1062.4 | 0.4 | 17 | 0 | All data processed in Iolite 3.6 | The distribution of Pb, U and Th in this standard appear to be more heterogeneous than in other standards we have analyzed (91500, Plesovice, Mud Tank). It is difficult to say based on the limited number of analyses we have done so far whether there is a significant correlation between Pb, Th, and U concentrations and measured age. We rejected 1 data point from RAK17 in this session, it gave older ages than the others and had low Pb concentration and high U/Th ration compared to the rest of the analyses.
G-Chron 2019 | A42 | LA-ICP-MS: QuadToF | 291.3 | 2.1 | 309 | 54 | 286 | 4.2 | 30 | 20 | 303-283 Ma | none | Single quadrupole | Agilent 7900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | 91500 | 1062Ma | ±10Ma |  | 2% | An Excel-based software ICPMSDataCal was used to perform off-line selection and integration of background and analyzed signals.
G-Chron 2019 | A44 | LA-ICP-MS: QuadToF | 293.2 | 1.8 | 288 | 37 |  |  | 50 | 47 | 307.8 - 275.8 | none | Single quadrupole | Agilent 7700x | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | 91500 | 1058.4 | 8.9 | 19 | 2 | NWR193UC with TruLine cell. 30s ablation. ICP-MS: TRA on 202Hg (0.02s), 204Pb (0.02s), 206Pb (0.04s), 207Pb (0.08s), 208Pb (0.04s), 232Th (0.02s), 238U (0.02s). 80 cycles for the age calculation. | One of Rak-17 fragments shows high concentration of U (2695 µg/g) and young 206Pb/238U age (144.1 Ma). It was rejected to the final data pool. Two analytical spots of Rak-17 contain high common Pb and low U concentrations and were rejected for the final data pool.
G-Chron 2019 | A46 | LA-ICP-MS: QuadToF | 305.7 | 1.8 | 417 | 51 | 309.3 | 5.1 | 24 | 22 | 312-290 | none | Single quadrupole | Agilent7900 | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | Yes | 91500 | 1062.5 Ma | 4.4 Ma | 10 | 0
G-Chron 2019 | A48 | LA-ICP-MS: QuadToF | 293.9 | 1 | 349 | 22 | 293.5 | 2.4 | 58 | 25 | 283.8-302.2 | none | Single quadrupole | Thermo iCAP Qc | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 13 | No | GJ-1 | 601.92 | 0.7 | 12 | 0 | 91500 (n=12), Mud Tank (n=12), and GZ7 (n=8) run and treated as unknowns during data reduction (iolite; Visual_AgeDRS). Concordia ages were respectively 1072.5±7.4 Ma, 723.9±12 Ma, and 525.5±2.9 Ma | We report here a weighted mean 206Pb/238U age in line with G-Chron requirements. However, we do not feel that a geologically meaningful central age exists for this material. Cathodoluminescence images show clear textural evidence for post-crystallisation alteration (11 out of 37 shards imaged). A small number of younger concordant (within our analytical resolution) single-analysis ages (n=6) were obtained, mostly from altered zones, which yielded a weighted mean age of 263.7 ± 4.3 Ma. One older concordant single-analysis age (325 ± 8 Ma) which fell outside the 2s dispersion of the dataset was also obtained. Trace element data obtained during the same analytical session (Ti, Lu, Hf) also indicate a complex history for this sample. While Hf content is relatively invariant (7841 ± 81 ppm, MSWD = 5.8, excluding two outliers), both Lu and Ti content are highly dispersed (ranges of 63-8 and 61-1 ppm, with MSWD of 696 and 204, respectively). Nominal temperatures obtained using the Ti-in-zircon thermometer range from 930 +12/-13 °C to 580 +29/-50 °C; the lower temperature values may reflect low Ti activity rather than true co-crystallisation with a Ti phase. As the briefing information supplied with the test material indicated sample collection from a pegmatitic schlieren hosted in a finer-grained magmatic rock, we tentatively interpret our c. 295 Ma 206Pb/238U age as recording crystallisation of the magmatic rock, and the c. 264 Ma age as recording formation of the pegmatite. The c. 325 older outlier may be inherited from the local country rock. We are willing to make all of our data available, and we encourage the organisers to contact us to request the full dataset.
G-Chron 2019 | A49 | LA-ICP-MS: QuadToF | 300.1 | 0.6 | 357.8 | 15.8 | 300.1 | 0.6 | 40 | 40 | 288-313 | none | Single quadrupole | Agilent 7900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | Yes | Yes | 5 | No | 91500 | 1062.4 | 3.5 | 40 | 0 |  | There is another group age of 292 Ma
G-Chron 2019 | A50 | LA-ICP-MS: QuadToF | 291.3 | 0.65 | 277 | 6 | 277.6 | 1.15 | 85 | 85 | 278.5 to 304.3 | none | Single quadrupole | Thermo IcapQ | Solid state 213 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | No | 91500 | 1065 Ma in the software, 1062 Ma observed | 0.28 | 22 | 0 |  | We also determine 207/235 which gave 291.6 Ma +/- 0.85 - it would have been nice to have that as a reporting option. We would not normally place much confidence in 207/206 for a zircon this young and 208/232 is really for information only. In some grains we noted slight evidence of alteration as evidence by a slightly enhanced abundance of La.
G-Chron 2019 | A51 | LA-ICP-MS: QuadToF | 293.4 | 0.4 | 385.3 | 39.4 |  |  | 48 | 42 | 283.1-299.8 | none | Single quadrupole | Thermo ICap Qc | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | 91500 | 1063 | 0.5 | 26 | 1 |  | We also have chemistry (trace and REE) available. Although not visible in age, the fragments are chemically zoned. Also, for laser ablation analyses, we don't normally report 207Pb/206Pb mean ages (unless samples older than 800 Ma).
G-Chron 2019 | A52 | LA-ICP-MS: QuadToF | 296.8 | 2.6 | 291 | 7.1 |  |  | 25 | 25 | 14 | none | Triple quadrupole | Agilent 8800 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 7 | Yes | 91500 | 1065 | 1 | 8 | 0 | Integration times 88Sr (0.01 ms), 238U (0.02 ms), 232Th (0.01 ms), 207Pb (0.09 ms), 206Pb (0.03 ms).
G-Chron 2019 | A58 | LA-ICP-MS: QuadToF | 294.2 | 0.7 | 304 | 18 | 292.6 | 2.3 | 25 | 25 | 13.6 | other | Single quadrupole | Agilent 7900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | 91500 | 1063.51 | 1.85 | 12 | 0
G-Chron 2019 | A59 | LA-ICP-MS: QuadToF | 285.6 | 0.7 | 301 | 6 |  |  | 42 | 41 | 294.3 - 277.4 | none | Single quadrupole | Agilent 7900 | Solid state 213 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 8 | Yes | Plesovice | 337 | 0.17 | 17 | 1
G-Chron 2019 | A61 | LA-ICP-MS: QuadToF | 295.47 | 0.71 | 309 | 16 | 293 | 2 | 44 | 38 | 300-291 | none | Single quadrupole | Agilent 7700x | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 40 to 90 µm | Spot |  | Helium | No | No | 5 | No | GJ-1 | 600.46 | 0.701 | 16 | 0 |  | Note: For common Pb correction, Anderson et al. (2002) correction was applied, but no correction was required.
G-Chron 2019 | A62 | LA-ICP-MS: QuadToF | 296.5 | 5.7 | 672.1 | 334.6 | 317.9 | 20.5 | 25 | 25 | 308-284 | none | Triple quadrupole | Agilent 7900 | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 6 | Yes | 91500 | 1064 |  | 8 | 0 | Off-line selection and integration of background and analyte signals, and time-drift correction and quantitative calibration for trace element analyses and U-Pb dating were performed by ICPMSDataCal.
G-Chron 2019 | A69 | LA-ICP-MS: QuadToF | 301.1 | 1.7 | 331 | 42 | 305.1 | 4.3 | 19 | 18 | 305-297 | none | Single quadrupole |  | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2
G-Chron 2019 | A70 | LA-ICP-MS: QuadToF | 293.9 | 1.9 | 309.3 | 42 | 292.5 | 5 | 27 | 22 | 304.2-281.5 | none | Single quadrupole | Agilent 7700X | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 6 | No | 91500 | 1062.4 | 6.6 | 8 | 0 | ASI RESOlution S155 and Agilent 7700x were used. Each analysis has 20s background, and 40 s sampe signal. Dwell times were 20 ms for 207Pb, 15 ms for 206Pb and 208Pb, 10 ms for 232Th and 238U. | 5 (out of 27) grains in this round of measurement do not have concordia U-Pb ages (concordance <95%). The contents of radiogenic Pb is low so that 207Pb/206Pb ratios are not in good precision. 176Hf/177Hf of RAK-17 zircon are 0.282826+/-0.000006 (1sigma, n=18).
G-Chron 2019 | A71 | LA-ICP-MS: QuadToF | 296.8 | 1.6 | 301.9 | 6.3 | 301.3 | 4.1 | 30 | 28 | 307.6 Ma - 291.9 M | 204Pb ancient | Single quadrupole | Thermo Fisher X Series 2 | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 6 | No | 91500 | 1064 Ma | 10 Ma | 12 | 0 |  | This test material is relatively uniform and can be used as a monitoring standard material during daily experimental testing.
G-Chron 2019 | A73 | LA-ICP-MS: QuadToF | 289.8 | 5.24 | 1.1 | 1.1 | 281.42 | 11.29 | 20 | 19 | 284.00-298.40 | other | Single quadrupole | Agilent 7900 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 40 to 90 µm | Line | 10 | Argon | Yes | No | 15 | No | 91500 | 1065 |  | 6 | 0 |  | ****************Enter Mean 207Pb/206Pb Age : NO DATA FOR THIS AGE, SORRY*************
Distribution | Lab Code | Analytical Technique | 206Pb/238U Age Result | 206Pb/238U Age Uncertainty | 207Pb/206Pb Age Result | 207Pb/206Pb Age Uncertainty | 208Pb/232Th Age Result | 208Pb/232Th Age Uncertainty | Total Measurements | Total Measurements in final pool | Total Range in final pool | Common Pb Correction | LA-ICP-MS: SFI / Which type of instrument did you use? | LA-ICP-MS: SFI / Please specify the make and model of the mass spectrometer that you used | LA-ICP-MS: SFI / Which laser wavelength did you use? | LA-ICP-MS: SFI / Which laser pulse length did you use? | LA-ICP-MS: SFI / Which laser fluence did you use? | LA-ICP-MS: SFI / What laser spot diameter did you use? | LA-ICP-MS: SFI / What sampling mode did you use? | LA-ICP-MS: SFI / Sampling mode - enter scan speed (μm/s) if applicable | LA-ICP-MS: SFI / Which carrier gas did you use? | LA-ICP-MS: SFI / Did you use a collision cell? | LA-ICP-MS: SFI / Did you use a reaction cell? | LA-ICP-MS: SFI / What laser repletion rate (Hz) did you use? | LA-ICP-MS: SFI / Did you use a signal smoothing device? | LA-ICP-MS: SFI / Which calibration material did you use? | LA-ICP-MS: SFI / What was the 206Pb/238U age (Ma) that you assigned to your calibration material? | LA-ICP-MS: SFI / What was the 206Pb/238U standard error (Ma) of the mean on your calibration material? | LA-ICP-MS: SFI / How many analyses were conducted on your calibration material? | LA-ICP-MS: SFI / How many determinations on your calibration material were rejected as outliers? | Other Relevant Information | Comments
G-Chron 2019 | A9 | LA-ICP-MS: SFI | 293.8 | 3.6 | 314.9 | 4.9 |  |  | 77 | 38 | 296.5 - 291.3 | none | Multi-collector sector field | Nu Plasma II | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | No | GJ-1 | 601.86 | 0.37 | 24 | 3 |  | Inhomogeneous grains (visible in CL) - analysis spots were discordant. Reported uncertainty is 2 sigma
G-Chron 2019 | A11 | LA-ICP-MS: SFI | 296.1 | 0.62 | 305.5 | 11.7 |  |  | 30 | 30 | 0.04552 - 0.04781 | none | Single collector sector field | Element II XR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 4 | No | 91500 | 1062.4 | 3.1 (2s) | 14 | 0
G-Chron 2019 | A13 | LA-ICP-MS: SFI | 295.51 | 0.44 | 295.8 | 3.7 |  |  | 73 | 73 | 17.9 | none | Single collector sector field | Nu Instruments Attom ES | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | No | 91500 | 1065 | 1.3 | 54 | 0 | Secondary Standard Temora2 gave a 206/238 age of 416.5+/-0.6 Ma | Pb-204 was not detected in the measured grains for this round.
G-Chron 2019 | A14 | LA-ICP-MS: SFI | 294.9 | 0.55 | 309.2 | 1.3 |  |  | 30 | 30 | 12.3 Ma | none | Multi-collector sector field | Nu Plasma I | Solid state 213 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 4 | No | GJ-1 | 605.95 | 0.5 | 6 | 0 |  | Most grains exhibited bright cathodoluminescence (CL) with corresponding low U and low radiogenic Pb. These 30 analyses are from two relatively dark CL grains. Pb loss is apparent. 11 total grains were mounted for analysis.
G-Chron 2019 | A15 | LA-ICP-MS: SFI | 299 | 2.6 | 326.9 | 7.4 | 292 | 11 | 68 | 33 | 284.4 - 311.5 | none | Multi-collector sector field | Nu Plasma II | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | Plesovice | 339.4 | 2.6 | 27 | 0 | Both GJ1 and 91500 were used as secondary standards. Iolite was used to reduce the data and correct for downhole fractionation. Isoplot was used to calculate the Concordia and weighted mean ages. | Data were corrected for 204Hg interference but not common Pb. Several grains exhibit inheritance of Archean age with a lower and upper intercept of 293.4 ± 3.5 and 2726 ± 380 Ma, respectively. At least one grain experienced extensive Pb loss.
G-Chron 2019 | A16 (*) | LA-ICP-MS: SFI | 294.3 | 3.2 | 323.7 | 67.3 |  |  | 314 | 314 | 312.4 - 282.9 | none | Single collector sector field | Thermo Element XR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot | 30s background followed by 60s measurement | Helium | No | No | 3 | Yes | GJ-1 | 601 Ma | SD = 3.55 Ma SEM = 0.25 Ma | 204 | 1 | Laser parameters were 30s backgroung followed by 60s of analysis laser firing at 33µm 3 Hz and 2J/cm2 | The uranium content is rather low: 85 ± 58 ppm (1s), similar to that of 91500 but very variable: 18 to 325 ppm. Unsurprisingly Th/U ratio is also variable: 0.58 ± 0.17 (1s). Reported in the Concordia diagram, it is obvious that few spots (about 3%) recorded the occurrence of common Pb. Age results are similar for tested thermally annealed grains and normal grains.
G-Chron 2019 | A30 | LA-ICP-MS: SFI | 297.7 | 4 | 293.8 | 6 |  |  | 600 | 450 | 310 - 290 | other | Single collector sector field | Thermo Scientific, Element 2 | Excimer 193 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot | 2 µm/s | Helium | No | No | 8 | Yes | Plesovice | 337.13 Ma | 3 % (10 Ma. aprox.) | 80 analysis | about 7 | Data reduction using Glitter | The evaluation of this Rack - 17, brought significant challenges due to the low concentrations of 207Pb, so we had to work at larger ablation sizes (50 um), in addition, we had to use the reference material 91500 and McClure as MR since they present concentrations similar, so these concentrations are in counting mode in the Element 2 detector, but it was difficult to normalize with plesovice since this is in analog mode due to the concentration of 238U.
G-Chron 2019 | A34 | LA-ICP-MS: SFI | 291.7 | 0.7 | 314 | 6.5 | 283.3 | 1.3 | 35 | 35 | 286 - 300 | none | Single collector sector field | Thermo Fisher Scientific Element XR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 8 | No | GJ-1 | 601.7 | 1 | 19 | 0 | tuning was carried out for maximum signal on 238U, lowest oxide (ThO<0.2 %) and Th/U close to one using line scan on NIST612 | no evidence for common Pb from 204 (Hg+Pb) signal
G-Chron 2019 | A43 | LA-ICP-MS: SFI | 288 | 4.8 | 308.8 | 5.1 |  |  | 34 | 34 | 38.9 | none | Single collector sector field | Element 2 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No |  | No | GJ-1 | 604.5 | 1.1 | 18 | 0
G-Chron 2019 | A45 | LA-ICP-MS: SFI | 295.3 | 1.9 | 310.7 | 11.7 | 297.7 | 3.7 | 25 | 24 | 301.2 - 290.4 | none | Single collector sector field | Thermo ElementXR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 4 | Yes | GJ-1 | 601.95 | 0.2 | 18 | 1 | Trace elements were measured "simultaneously" with the dating. | The 206/238 age was corrected for differences in dose and Th disequilibrium (Sliwinski et al. 2017) The 206/238 age uncertainty contains a measured long term excess variance of 1.2% (2s) The 208/232 age uncertainty contains an estimated extra variance of 2% (2s) The 207/206 age ubcertainty contains an estimated extra variance of 2% (2s) I would not use the 207/206 age for Zircons in this age range, as it is very sesitive to common Pb, and no common Pb correction is applied. One laser single point analysis was high in U >1500 ppm and likely sufferd high Pb loss.
G-Chron 2019 | A53 | LA-ICP-MS: SFI | 295.5 | 2 | 305 | 33.6 |  |  | 6 | 6 | 292.8 - 297.5 | none | Single collector sector field | ELEMENT XR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Line | 2 µm/s | Helium | No | No | 10 | No | GJ-1 | 600.62 | 2.4 | 5 | 0
G-Chron 2019 | A54 | LA-ICP-MS: SFI | 292.4 | 0.48 | 740 | 25.5 | 262.2 | 1.8 | 30 | 27 | 296.4-289.4 | none | Multi-collector sector field | Thermo Neptune Plus | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 7 | Yes | 91500 | 1062 | 4.7 Ma | 18 | 0 | 2 J/cm2 fluence, 35 um spot, all masses measured on Faraday cups, ablated aerosol split between MC-ICP and Quad ICP, for simultaneous trace element analysis. | Has CL zoning and REE pattern consistent with igneous crystallization, with positive Ce/Ce* (mean = 41) and negative Eu/Eu* (mean = 0.2); however is significant variation in MREE concentration from grain to grain (Yb/Gd range = 8-32). Th/U also vary considerably from grain to grain, with range = 0.35-1.0. High Th/U grains correlate with lower Ce/Ce* and Yb/Gd, and higher Nb/Ta. U-Pb data are almost perfectly concordant, but include a few older regions yielding 206/238 dates of 300-316 Ma (isotopic inheritance?). Excluding these three spots, Concordia Age = 292.41 +/- 0.91 Ma (MSWD =n 0.64). Given it's apparent Paleozoic age and the internal precision and reproducibility on the measured ratios, I consider the 206Pb/238U Age and Concordia Age to be the most robust ages for this zircon
G-Chron 2019 | A63 | LA-ICP-MS: SFI | 295.8 | 1.8 | 302.6 | 3.5 |  |  | 17 | 17 | 285.7 - 300.0 | none | Multi-collector sector field | Nu Plasma II | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 7 | No | 91500 | 1065.4 | 0.3 | 22 | 4 |  | Below are given concordia ages defined by IsoplotR (Vermeesh et al., 2018) using Com Pb correction by Stacey and Kramers Primary Standard 91500 = 1063.84 +- 0.73; n= 17 Secondary Standard Temora = 416.14 +- 0.81; n = 8 Plesovice treated as unknown = 336.94 +- 0.46; n = 14 G-Chron = 296.63 +- 0.41; n = 17
G-Chron 2019 | A65 | LA-ICP-MS: SFI | 295 | 3 | 299.9 | 28 | 292.6 | 15.9 | 24 | 24 | 13.9 | none | Single collector sector field | Element XR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | No | 91500 | 1061 | 1 | 24 | 0 |  | U contents vary from 39 ppm to 388 ppm; Th contents vary from 22 to 238 ppm; U/Th is around 1.70
G-Chron 2019 | A68 | LA-ICP-MS: SFI | 296.2 | 4.6 | 294.1 | 57.3 | 292 | 13.8 | 36 | 35 | 313.3 - 290.9 | none | Single collector sector field | Element XR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | No | 91500 | 1062.4 | 2 | 16 | 0 |  | Some of the grains seem to be slightly older than the average.
G-Chron 2019 | A72 | LA-ICP-MS: SFI | 293.6 | 3.7 | 305 | 25.1 | 280.3 | 9.1 | 120 | 72 | 11 | other | Single collector sector field | Thermo-Element2 | Solid state 213 nm | Nanosecond | > 5 to 10 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 5 | Yes | GJ-1 | 600.4 | 6.9 | 95 | 29 | Run: 4 GJ-1, 91500, Plesovice, 10 unknowns, 4 GJ-1. Glitter 4.2.2. Each run was processed separately. The relative uncertainty (1s) of the ratio 207/206 should not exceed 10% of the average. | At the preliminary stage of research, 18 grains of test material were analyzed. Of this amount, 5 grains were selected for subsequent analysis. The analyzed 120 dots were made on the selected 5 grains.The selection of 5 grains for analysis was performed according to the criterion: the relative uncertainty (1s) of the isotop ratio 207Pb/206Pb should not exceed 10% of the average.
G-Chron 2019 | A74 | LA-ICP-MS: SFI | 297.1 | 3.5 | 310 | 27.6 | 297.3 | 6.9 | 30 | 29 | 291.7 - 304.9 | none | Single collector sector field | Thermo Fischer Scientific Element 2 | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | Yes | 91500 | 1062 | 0 | 30 | 0
G-Chron 2019 | A75 | LA-ICP-MS: SFI | 292 | 5.5 | 0 | 0 |  |  | 56 | 56 | 305.9 - 277 | none | Single collector sector field | Thermo Scientific - Element XR | Excimer 193 nm | Nanosecond | <= 5 J cm-2 | > 10 to 40 µm | Spot |  | Helium | No | No | 10 | No | 91500 | 1063 | 0.39 | 22 | 0 | Iolite 3 used for data reduction. | We report 207Pb/206Pb Age as 0.0 ± 0.0 because using our routine procedure, we aren´t able to measure 207Pb signal with reasonable quality. It was near our background.
| (*) Laboratory method orginally missclassified as "other" |  |  |  | Artificial input to accommodate system restrictions | [5 empty] | Suspected input error